Universal Single-Use Bronchoscope Tracker

Updated Engineering Plan and Design Blueprint v4

June 30, 2026 | Nonclinical simulation prototype

Figure 1. v4 blueprint showing the wired architecture, clamshell Module B, wiper zone, V-groove tracking saddle, and differential I2C cable.

# 0. What changed in v4

| Area | Previous direction | v4 decision |
| --- | --- | --- |
| Module B lubricant handling | Rely on SQUAL warning and cleaning | Add replaceable proximal felt/foam wiper ring before photogate and sensor. |
| Module B standoff | Snug bore plus leaf spring | Use diameter-specific V-groove tracking saddle around the optical aperture with compliant opposing cap. |
| Module A cable | Direct single-ended I2C over ~1.2 m, PCA9615 fallback | Make PCA9615 differential I2C the default wired architecture. |
| Wireless lever module | Possible Phase 4 BLE option | Remove from roadmap; keep all sensing wired and synchronized. |
| Compatibility language | Any disposable bronchoscope | Adapter-supported disposable bronchoscopes, validated model by model. |
| Module B geometry | Pass-through bore | Split clamshell tracker with swappable insert cassette, wiper cassette, and flange. |

# 1. One-sentence concept

A two-module, clip-on, open-hardware training controller that lets a disposable bronchoscope act as a three-degree-of-freedom simulator input device without modifying the scope: Module A tracks thumb-lever flexion and Module B tracks insertion depth and roll at the cord.

# 2. Realistic development target

The realistic v1 target is a reliable adapter-supported bronchoscopy simulator controller, first validated on one disposable bronchoscope family, then expanded through parametric adapters. The project should not claim true universality until each scope family has a validated adapter, insert, and calibration profile.

# 3. Non-negotiable design rules

1.    Training-only use: simulation, education, and bench validation only; not for patient care or clinical navigation.

2.    No scope modification: no adhesive, drilling, magnets attached to the scope, disassembly, or interference with real suction/working-channel controls.

3.    Adapter-supported compatibility: v1 is validated on selected disposable bronchoscopes; additional scopes require published adapter profiles and repeat validation.

4.    Wired baseline: the lever module connects through a flexible cable using differential I2C.

5.    Lubricant-tolerant optical zone: include a replaceable proximal wiper and SQUAL-based quality gate.

6.    Controlled standoff: center the cord with a V-groove saddle and compliant cap, not a loose bore.

7.    Browser-native runtime: USB HID gamepad for runtime, CDC/Web Serial or CLI for calibration.

8.    Objective phase gates: no multi-scope expansion until tracking, standoff, friction, and cable stability pass bench testing.

# 4. System architecture

MODULE A - LEVER MODULE      -> 1.2 m differential I2C cable -> MODULE B - CORD TRACKER
C-clamp + AS5600 + buttons       PCA9615 on both ends           wiper + gate + V-groove + optical sensor + RP2040 USB

| Degree of freedom | Sensed by | Location | Principle |
| --- | --- | --- | --- |
| Tip flexion | AS5600 magnetic encoder | Module A hinge | Follower-arm angle maps to thumb-lever angle. |
| Insertion depth | Optical tracking sensor | Module B cord saddle | Optical Y counts after 2x2 calibration matrix. |
| Roll | Optical tracking sensor | Module B cord saddle | Optical X counts after 2x2 calibration matrix. |
| Buttons | Tactile switches | Module A body | Simulation controls; real suction button untouched. |

All high-rate data are fused in Module B firmware and emitted as one synchronized USB HID report. This avoids browser-side timing mismatch between independent BLE and USB streams.

# 5. Module A - lever module blueprint

## 5.1 Mechanical design

Module A is a printed clamp that attaches to the bronchoscope handle neck without altering the scope. It carries its own hinge and follower arm. The follower arm touches the thumb lever through a small low-friction roller or rounded shoe.

Two-piece C-clamp with TPU or silicone pads.

Captive M3 thumb screw or over-center latch.

Hard datum surface for repeatable clamp placement.

Follower arm pivot on a steel or shoulder-screw axle.

Light torsion spring to maintain lever contact.

Roller or polished rounded shoe with 1-2 mm compliance.

Optional simulator buttons on the module body.

Cable strain relief directed toward the scope umbilical/cord path.

## 5.2 Sensor and calibration

The follower-arm hinge uses an AS5600 magnetic rotary sensor and a diametric magnet embedded in the follower hub. Calibration captures neutral, full up, full down, and an optional midpoint check to detect shoe slippage or clamp misplacement.

## 5.3 Module A electronics

| Subcomponent | Purpose |
| --- | --- |
| AS5600 | Lever angle measurement. |
| Diametric magnet | Rotates with follower arm. |
| PCA9615-A | Converts local single-ended I2C to differential I2C for cable. |
| MCP23008/PCF8574 or similar | Reads optional buttons over the same local I2C segment. |
| ESD/TVS protection | Cable robustness. |
| JST-GH or locking connector | Strain-resistant detachable cable. |

# 6. Module B - cord tracker blueprint

Module B is the critical module. v4 changes it from a simple pass-through bore into a split clamshell tracker with controlled lubrication handling, focal distance, and cable management.

## 6.1 Zone layout

1.    Wiper zone: replaceable slotted felt or high-density foam ring to remove silicone lubricant, condensation, and phantom residue before the optical window.

2.    Zero/gate zone: IR photogate or optical interrupter for tip passage/zero sanity check; manual zero at the flange plane remains available.

3.    Tracking zone: diameter-specific insert cassette with V-groove saddle around the optical window, compliant top cap, and optical sensor below.

4.    Mounting flange: standard interface to phantom inlet, manikin adapter, or desk stand.

## 6.2 Lubrication barrier

The wiper is now a baseline component. The tracker must maintain adequate SQUAL after a realistic silicone-lubricant challenge; a software warning alone is insufficient.

Tool-free sliding cassette holding a replaceable C-shaped felt or foam ring.

Slotted geometry so it can be replaced without threading over the distal tip.

Ring ID approximately cord OD minus 0.1 to 0.2 mm for gentle contact, tuned by material.

Thickness 3-5 mm in the direction of travel.

Spare wipers stored in the kit.

Firmware counter reminding users to replace the wiper after a defined number of sessions or if SQUAL falls.

## 6.3 V-groove standoff control

The insert should use two polished rails forming a shallow V around the optical aperture. The cord sits in the V, while a compliant opposing cap supplies gentle downward force. The V-groove should be interrupted around the optical aperture rather than forming an opaque floor.

| Parameter | Starting value |
| --- | --- |
| Insert family | One cassette per cord OD range. |
| Cord radial clearance | Cord OD + 0.2 to 0.4 mm outside the tracking saddle. |
| V-groove angle | 90-120 degrees, test both. |
| Optical slot width | 1.5-2.5 mm, sensor dependent. |
| Compliant cap travel | 0.5-1.5 mm. |
| Added insertion force | Target <0.5 N after wiper + V-groove are installed. |
| Insert material | PETG/ASA body; polished PTFE/POM or resin insert surface if friction is high. |

## 6.4 Clamshell and insert strategy

Hinged or fully separable two-piece body.

Captive latch or thumb screw.

Swappable insert cassette for each cord diameter family.

Swappable wiper cassette independent of the insert cassette.

Hard stop defining the flange/zero plane.

Cable exit and USB port protected from hand contact and phantom moisture.

Drainage or removable cover so residue does not pool over the optical sensor.

# 7. Optical sensor strategy

The PAT9125 remains the preferred clean open-hardware candidate if it passes Phase 0. v4 treats optical tracking as environment-dependent rather than guaranteed.

| Path | When to use | Comment |
| --- | --- | --- |
| PAT9125 single sensor | Default if SQUAL and accuracy pass wet/dry tests. | Lowest cost and simplest open design. |
| PMW3360 or similar | Use if PAT9125 fails on cord finish or speed. | Better performance but more firmware/licensing complexity. |
| Dual optical sensors 90 degrees apart | Use if single-sensor cross-coupling remains high. | Adds cost and alignment burden. |
| Wheel encoder cassette | Use if optical tracking fails after lubrication mitigation. | Less elegant, but robust and publishable fallback. |

Firmware should continuously track SQUAL or equivalent surface-quality metric, delta count saturation, dropout frequency, unexpected jumps, and calibration-motion correlation.

# 8. Electronics and signal integrity

## 8.1 Default wired architecture

Single-ended I2C over a moving 1.2 m cable is no longer the baseline. The AS5600 and button expander communicate locally to PCA9615-A in Module A. The cable carries differential I2C to PCA9615-B in Module B, which connects to the RP2040.

AS5600 + buttons -> local I2C -> PCA9615-A -> differential cable -> PCA9615-B -> RP2040

| Conductors | Function |
| --- | --- |
| Twisted pair 1 | SDA+ / SDA-. |
| Twisted pair 2 | SCL+ / SCL-. |
| Power | 3.3 V or 5 V, selected by PCB design. |
| Ground/shield | Ground reference and shield termination. |

## 8.2 Module B main PCB

| Component | Purpose |
| --- | --- |
| RP2040/Pi Pico or custom RP2040 board | Sensor fusion, USB HID, CDC serial. |
| PCA9615-B | Differential I2C receiver/transceiver for Module A. |
| PAT9125 or fallback optical sensor | Cord motion tracking. |
| IR photogate | Tip passage/zero sanity check. |
| Calibrate button | Local zero/calibration mode. |
| RGB LED | Tracking OK, low SQUAL, cable fault, calibration. |
| USB-C | Power and data. |
| Optional EEPROM | Hardware revision and calibration defaults. |

# 9. Firmware design

Optical sensor read: 500-1000 Hz target.

Lever sensor read: 250-500 Hz target over differential I2C.

Fusion/report loop: 125-250 Hz HID report rate.

CDC serial: raw counts, SQUAL, matrix values, diagnostics.

## 9.1 Calibration model

[ depth_mm ]   [ a  b ] [ dy_counts ]
[ roll_deg ] = [ c  d ] [ dx_counts ]

1.    Pure 100 mm insertion using ruler jig.

2.    Pure 360 degree roll using roll jig.

3.    Combined helical motion to check cross-coupling.

4.    Lubricated-cord repeat after wiper pass.

## 9.2 HID report design

| HID axis | Meaning |
| --- | --- |
| Axis 0 | Flexion, -1 to +1. |
| Axis 1 | Depth normalized to selected airway path length. |
| Axis 2 | sin(roll). |
| Axis 3 | cos(roll). |
| Buttons | Simulator action buttons. |

## 9.3 Fault handling

Low SQUAL: freeze optical accumulation only if dropouts exceed threshold; otherwise continue with warning.

Cable fault: reinitialize lever I2C bus and log event.

Stuck I2C bus: firmware attempts bus recovery, then requests reconnect if needed.

Calibration mismatch: reject profile if helical validation exceeds error threshold.

# 10. Web app integration

Runtime remains USB HID/Gamepad. Calibration uses Web Serial on supported desktop browsers, with a CLI fallback for locked-down systems. v4 removes Web Bluetooth. The wired lever cable mimics a bronchoscope umbilical, avoids battery management, avoids browser BLE compatibility problems, and keeps all sensor fusion under one timestamped firmware loop.

| Component | Purpose |
| --- | --- |
| ScopeInputProvider | Abstracts keyboard, gamepad, and hardware scope input. |
| /hardware setup route | Connect, live bars, SQUAL display, profile selection, calibration. |
| useScopeInput() hook | Client-only input hook for simulation runtime. |
| Local profile store | Insert ID, lever calibration, 2x2 matrix, SQUAL thresholds. |
| Optional cloud profile | Later Supabase storage for classroom kits and competency dashboards. |

# 11. Build phases and exit criteria

| Phase | Work | Exit criterion |
| --- | --- | --- |
| 0 - Bench de-risking | Dry/lubricated cord, wiper, V-groove, lateral torque, helical motion, SQUAL logging. | Depth <= +/-5 mm over 300 mm; roll <= +/-10-15 deg; dropout <1%; lube passes after wiper. |
| 1 - Module B alpha | Clamshell with wiper cassette, V-groove insert, optical board, RP2040. | Insertion and roll drive debug visualizer after wet/dry tests. |
| 2 - Module A alpha | Lever clamp with AS5600, button expander, PCA9615-A. | Stable lever signal during 30-min moving-cable test; no bus lockups. |
| 3 - Integrated HID prototype | Fuse all data into HID gamepad report with calibration UI. | All 3 DOFs drive simulator with perceived latency <30 ms. |
| 4 - Robustness/usability | Attach/detach cycles, wiper replacement cycles, fellow-facing usability. | 50-100 attachment cycles; setup under 2 min; no structural failure. |
| 5 - Adapters/publication | Additional adapters and manuscript package. | Bench validation + face/content validity + reproducible files. |

# 12. Validation plan

| Test | Method | Output |
| --- | --- | --- |
| Linear depth | 300 mm ruler/linear stage | Error, repeatability, drift. |
| Roll | Rotary jig, 360/720 degrees | Roll error and wrap behavior. |
| Helical motion | Simultaneous insertion + roll | Cross-coupling residual after matrix calibration. |
| Lubrication challenge | Silicone lube and phantom residue | SQUAL, error, dropout with/without wiper. |
| Friction | Spring scale or force gauge | Added insertion force. |
| Cable integrity | Moving-cable stress test | Dropped frames, I2C recovery events. |
| Attach/detach | Repeated mounting cycles | Calibration drift, mechanical wear. |

Simulation validity should compare novice, fellow, and attending performance on named-segment navigation tasks, including time, path efficiency, collision rate, branch-selection accuracy, and user-rated realism.

# 13. Risk register v4

| Risk | Current likelihood | Mitigation now in baseline |
| --- | --- | --- |
| Lubricant/residue blocks optical tracking | High | Replaceable wiper before photogate/sensor; SQUAL gate; wet testing in Phase 0. |
| Cord wanders away from focal plane | High | V-groove saddle around aperture with compliant opposing cap. |
| Optical tracking unreliable on specific cord finish | Medium-high | Test PAT9125 and fallback sensor; wheel-encoder cassette branch. |
| I2C lever bus unstable over cable | Medium without extender; low with extender | PCA9615 default on both ends; moving-cable stress test. |
| BLE latency/sync problems | Removed | No wireless phase. |
| Added friction changes scope feel | Medium | Wiper material testing; V-groove polish/PTFE; force threshold. |
| Clamp fit varies across scopes | Medium | Parametric adapter library and per-model validation. |
| Classroom durability | Medium-high | Heat-set inserts, captive fasteners, spare wipers, replaceable insert cassettes. |

# 14. Bill of materials - realistic v4 estimate

| Subassembly | Expected prototype cost | At-scale kit target |
| --- | --- | --- |
| Module A mechanical parts | $10-20 | $5-10 |
| AS5600, magnet, button expander | $5-8 | $3-5 |
| Module A PCA9615/PCB/connectors | $10-20 | $5-10 |
| Differential cable and strain relief | $8-15 | $4-8 |
| Module B RP2040/USB/status electronics | $10-20 | $6-12 |
| Optical sensor path | $8-20 | $5-12 |
| Module B PCA9615/PCB/connectors | $10-20 | $5-10 |
| Photogate and wiper cassette materials | $5-15 | $2-6 |
| Printed housings/inserts/hardware | $15-30 | $8-15 |
| Total | $81-168 | $43-88 |

Planning number: $100-180 for early prototypes and $70-100 for a polished educational kit after PCB consolidation and print optimization.

# 15. CAD and repo package

/hardware
  /cad
    module_a_lever_clamp.build123d.py
    module_b_clamshell_tracker.build123d.py
    insert_profiles.yaml
    flange_standard.step
  /stl
    module_a_base/ module_a_shoes/ module_b_body/ v_groove_inserts/ wiper_cassettes/ desk_stand/
  /electronics
    module_a_pca9615_as5600/ module_b_rp2040_tracker/ wiring_diagrams/ bom.csv
  /firmware
    platformio/ drivers/ calibration/ releases/uf2/
  /web
    hardware_setup_route/ useScopeInput/
  /docs
    build_guide.md calibration_guide.md validation_protocol.md scope_profile_contribution_guide.md

# 16. Immediate next actions

1.    Build a benchtop V-groove optical jig before printing a full enclosure.

2.    Test wiper materials using dry, lubricated, and phantom-residue cord conditions.

3.    Order two PCA9615 breakout boards or design the small Module A/B PCBs immediately.

4.    Remove BLE tasks from the roadmap and issue list.

5.    Define the first validated scope family and insert dimensions.

6.    Create a validation spreadsheet before running tests so failures become design data rather than anecdotes.

# 17. Source notes

User-supplied v2 engineering plan: universal-scope-tracker-plan-v2(1).md.

ams OSRAM AS5600 product page: https://ams-osram.com/products/sensor-solutions/position-sensors/ams-as5600-position-sensor

PixArt PAT9125EL public references: https://www.pixart.com/search/%26page%3D1%26keyword%3DPAT9125EL and https://www.epsglobal.com/Media-Library/EPSGlobal/Products/files/pixart/PAT9125EL-TKITPAT9125EL-TKMT.pdf

NXP PCA9615 product datasheet: https://www.nxp.com/docs/en/data-sheet/PCA9615.pdf

NXP I2C-bus specification/user manual: https://www.nxp.com/documents/user_manual/UM10204.pdf

MDN Web Bluetooth API: https://developer.mozilla.org/en-US/docs/Web/API/Web_Bluetooth_API

MDN Gamepad API: https://developer.mozilla.org/en-US/docs/Web/API/Gamepad_API

MDN Web Serial API: https://developer.mozilla.org/en-US/docs/Web/API/Web_Serial_API